ADS1299 — Analog Front End IC Texas Instruments │ In-Ear EEG Wearable for Epileptic Seizure Pre-Ictal Detection

Overview

The ADS1299 is a medical-grade, 8-channel, 24-bit simultaneous-sampling biopotential AFE by Texas Instruments, purpose-engineered for EEG, ECG, and EMG acquisition. It consolidates the complete analog signal chain — instrumentation amplification, programmable gain, delta-sigma ADC conversion, active noise cancellation, lead-off detection, and internal reference routing — into a single 64-pin TQFP IC, requiring no external op-amps, clamp circuits, or signal conditioning components.

Electrical Specifications

| Parameter | Value | Design Relevance |
| --- | --- | --- |
| ADC Architecture | Delta-sigma, 24-bit | Oversampling + decimation — inherent noise shaping, no external anti-alias filter needed |
| Resolution | 24-bit (8,388,607 codes) | LSB = 0.022 µV at Gain=24, VREF=4.5 V — resolves sub-microvolt pre-ictal brainwave shifts |
| Sampling Method | 8 simultaneous ADC cores | All 5 electrodes sampled at identical microsecond — zero inter-channel phase lag |
| Data Rate | 250 / 500 / 1k – 16k SPS | 250 SPS selected — Nyquist at 125 Hz covers full EEG spectrum (0.5–100 Hz) |
| PGA Gain | 1×, 2×, 4×, 6×, 8×, 12×, 24× | Gain=24 amplifies 5–100 µV EEG to full ADC input range |
| Input-Referred Noise | ~1.0 µ Vpp at 250 SPS | Below EEG signal floor — preserves signal integrity for feature extraction |
| CMRR | >80 dB typical | Rejects common-mode EMG, motion, and EMI noise across electrode pairs |
| Input Impedance | >1 GΩ | Prevents loading of high-impedance Ag/AgCl electrode interface |
| Internal Reference | 2.4 V or 4.5 V selectable | 4.5 V selected — maximises dynamic range at Gain=24 |
| Supply Current | ~5 mA active (all 8 channels) | Compatible with LiPo-powered wearable — full-day battery feasible |
| Supply Voltage | 2.7 – 5.5 V | 3.3 V ESP32 logic compatible — no level shifting required |
| SPI Interface | Up to 20 MHz, Mode 1 | 27-byte frame read in 54 µs at 4 MHz — host MCU idle 98.6% per sample cycle |

Key Integrated Features

▸ Simultaneous Sampling — 8 independent ADC cores capture all electrodes at the identical microsecond. Critical for pre-ictal inter-channel phase synchronisation analysis — multiplexed ADCs introduce up to 4 ms skew that corrupts this measurement.

▸ BIAS Drive Amplifier — Senses common-mode interference across electrodes, inverts it, and drives it back into the body via a dedicated electrode. Actively cancels 50/60 Hz powerline noise before ADC conversion — equivalent to a clinical driven-right-leg circuit.

▸ SRB1 Reference Bus — Single mastoid reference electrode connects to SRB1 pin; ADS1299 internally routes this to the negative input of all 8 channels. Eliminates complex PCB routing and prevents the reference wire acting as a noise antenna.

▸ Lead-Off Detection — Injects a small AC/DC current per channel. Impedance change from electrode lift flags a status bit in every 27-byte data frame. Firmware discards corrupt samples and alerts the user instantly.

▸ Input Clamp Diodes — Built-in ESD protection diodes on every channel input — no external TVS diodes, clamp ICs, or protection resistors required on the PCB.

Ag/AgCl — Silver / Silver Chloride Electrode Electrode Material Justification │ In-Ear EEG Wearable for Epileptic Seizure Pre-Ictal Detection

Overview

Ag/AgCl (silver/silver chloride) is the universal standard for clinical biopotential electrodes, used in every hospital-grade EEG, ECG, and EMG system worldwide. Its selection over gold, graphene, and platinum is based on its unique classification as a non-polarizable electrode — the single material property that directly determines signal quality in the 0.5–100 Hz EEG acquisition band.

Electrochemical Justification

| Electrode Type | Examples | Interface Behaviour | EEG Signal Impact |
| --- | --- | --- | --- |
| Polarizable | Gold, Graphene, Platinum | Charge accumulates at skin interface — does not transfer freely | Baseline drift + elevated 1/f noise in 0.5–30 Hz band |
| Non-Polarizable ✓ Selected | Ag/AgCl | Free charge transfer via reversible Cl⁻ redox reaction | Stable baseline, flat noise floor across full EEG spectrum |

Governing reaction:  Ag + Cl⁻  ⇌  AgCl + e⁻     Half-cell potential: +0.197 V (stable, temperature-independent)

The reversible Cl⁻ redox reaction maintains a fixed half-cell potential regardless of sweat composition, temperature, or duration. This eliminates the slow voltage wander seen with polarizable electrodes — a critical requirement as the seizure prediction algorithm cannot distinguish electrode drift from genuine pre-ictal slow-wave activity.

Performance at Mastoid Placement

▸ Self-activating impedance reduction — Sweat at the mastoid region (20–100 mM NaCl) reacts with the silver surface to continuously rebuild the AgCl layer. Electrode-skin impedance drops from ~100–200 kΩ on dry contact to ~1–5 kΩ within 15 minutes — no conductive gel or skin preparation required.

▸ Continuous self-healing — Sweat Cl⁻ ions replenish the AgCl layer as it degrades through ammonia attack or mechanical abrasion, extending functional electrode life beyond passive metal electrodes of equivalent thickness.

▸ Antimicrobial protection — Trace Ag⁺ ions released at the contact surface (Ksp = 1.8 × 10⁻¹⁰) are bactericidal and fungistatic at sub-toxic concentrations — preventing microbial growth at the behind-ear skin contact zone during 24/7 wear.

Electrode Technical Specifications

| Parameter | Value | Notes |
| --- | --- | --- |
| Half-cell potential | +0.197 V | Stable across body temperature (35–37°C) and sweat pH (4.5–7.5) |
| Skin impedance (dry) | 100–200 kΩ | Initial contact before sweat activation — ADS1299 >1 GΩ input prevents loading |
| Skin impedance (wet) | 1–5 kΩ | After 10–15 min wear — sweat Cl⁻ self-activates AgCl layer |
| Noise floor (0.5–30 Hz) | <0.5 µV RMS | Flat 1/f noise profile across full EEG band — superior to all polarizable alternatives |
| Electrode substrate | Medical PDMS (Sylgard 184) | ISO 10993 biocompatible, Shore A hardness 43 — soft earbud-like contact |
| AgCl layer thickness | 5–20 µm | Formed by 45-min NaOCl chlorination of 99.9% Ag wire (0.5 mm dia) |
| Protective overcoat | Parylene-C, 2–5 µm | Vapor-deposited — blocks UV, H₂O₂, NH₃, SLS while retaining Cl⁻ micropore access |
| Operating temperature | 15–45°C | Stable across full ambient + body temperature range for wearable use |
| Biocompatibility | ISO 10993 certified | 60+ years clinical use — no skin toxicity, no allergic sensitisation (except ~1–3% silver allergy) |
| Electrode lifespan | 18–24 months (Parylene-C coat) | Uncoated ~6 months. Replaceable earbud tip design enables low-cost renewal |
| Cost (5 electrodes) | ₹110–160 DIY | 99.9% Ag wire + NaOCl chlorination — no specialised equipment required |

Material Comparison

| Material | Type | 1/f Noise | Durability | Cost (5 el.) | Decision |
| --- | --- | --- | --- | --- | --- |
| Ag/AgCl | Non-polarizable | Lowest | 18–24 months* | ₹110–160 DIY | ✓ Selected |
| Gold | Polarizable | Higher | 5+ years | Moderate (sputtered) | × Noise |
| Graphene | Polarizable | Higher | Variable | High (complex fab) | × Noise |
| Platinum | Polarizable | Higher | 5+ years | High (sputtered) | × Noise |
| IrOx | Pseudo non-pol. | Low | 5–10 years | ₹50,000+ (PVD) | Future upgrade |

* With Parylene-C overcoat. Uncoated lifespan: ~6 months. Replaceable earbud tip design allows low-cost periodic renewal.