AgCl

What it is

- Silver base with a silver chloride (AgCl) surface layer

- Non-polarizable electrode — charge transfers freely, doesn't build up

- Every clinical EEG, ECG, EMG device uses this material

Why it works for EEG

- Lowest 1/f noise of any electrode material in 0.5–100Hz range

- Stable half-cell potential — baseline doesn't drift

- Self-activating in ear canal — sweat Cl⁻ ions react with AgCl layer, impedance drops automatically over 10–15 mins

- ADS1299 is literally designed assuming Ag/AgCl on input pins

Physical properties

- Lightweight as thin film on silicone

- Flexible when used as ink or coating

- Grey-purple colour when properly chlorinated

- Density 10.5 g/cm³ (bulk), negligible as thin film

Biocompatibility

- 60+ years of clinical use, billions of patients

- ISO 10993 certified for skin contact

- Only risk: 1–3% silver allergy — do patch test first

- No toxicity, no skin damage under normal use

Degradation

- AgCl layer depletes over 6–12 months daily use

- Degrades faster under UV light — protected inside ear canal

- Can be re-chlorinated with FeCl₃ solution (resets electrode)

- Design earbud tip as replaceable module

Reactions

1. Photodecomposition

- AgCl + light → Ag + ½Cl₂

- Electrode efficiency problem AgCl converts back to metallic silver. Electrode shifts from non-polarizable to polarizable behaviour. Baseline starts drifting continuously. 1/f noise increases. ADS1299 sees a slowly wandering DC offset that your algorithm cannot distinguish from actual pre-seizure brain activity.

- Prevention — Parylene-C UV-blocking overcoat

| Advantages | Disadvantages |
| --- | --- |
| Blocks UV/visible light from reaching AgCl | Adds fabrication step, increases cost |
| Extends electrode lifespan significantly | Micropores can clog with sebum over time |
| Maintains non-polarizable behaviour longer | Very slightly increases electrode-skin distance |
| Transparent — no aesthetic change | Requires vapor deposition equipment |

2. Silver tarnishing

- 2Ag + H₂S → Ag₂S + H₂

- 2Ag + S compounds → Ag₂S

- Electrode efficiency problem Ag₂S is a semiconductor, not a conductor. Forms as black patches across electrode surface. Creates uneven conductivity — some zones of electrode read signal, others don't. Effective electrode area shrinks progressively. Signal amplitude drops, noise increases.

- Prevention — Anti-tarnish polymer coating

| Advantages | Disadvantages |
| --- | --- |
| Physically blocks H₂S and sulfur compounds | Adds thin insulating layer, raises baseline impedance slightly |
| Cheap and easy to apply | Doesn't help during storage without sealed case |
| No impact on electrode shape | Needs reapplication after mechanical wear |

- Prevention — Sealed storage case when not in use

| Advantages | Disadvantages |
| --- | --- |
| Eliminates air exposure completely | Only works when not worn — zero help during daily use |
| Zero impact on electrode chemistry or impedance | Adds inconvenience to user routine |
| Cheap, no fabrication change | — |

3. Ammonia attack from sweat

- AgCl + 2NH₃ → [Ag(NH₃)₂]⁺ + Cl⁻

- Electrode efficiency problem AgCl layer physically dissolves into solution. Layer thickness reduces month by month. Electrode progressively loses non-polarizable behaviour. Baseline drift worsens gradually — not suddenly. Hard to detect until signal quality is already severely compromised. Primary cause of lifespan limitation.

- Prevention — Parylene-C microporous overcoat

| Advantages | Disadvantages |
| --- | --- |
| Slows ammonia penetration significantly | Doesn't stop it completely — only slows |
| Sweat Cl⁻ still passes through micropores for self-regeneration | Micropores add complexity during fabrication |
| Single coating protects against multiple reactions | Effectiveness reduces as micropores clog |

- Prevention — Replaceable earbud tip design

| Advantages | Disadvantages |
| --- | --- |
| Fresh electrode fully restores performance | User must remember to replace every 6–12 months |
| No chemistry change needed | Recurring cost for user |
| Clean business model for product | Replacement tip must be biocompatible and standardised |

4. Hair product / bleach attack

- Ag + H₂O₂ → Ag⁺ + OH⁻ + H₂O

- Electrode efficiency problem Most aggressive reaction. H₂O₂ directly oxidises the silver base, not just the AgCl layer. Damage is irreversible — silver base cannot be restored. Creates permanent dead zones on electrode. Signal loss in those areas is permanent. Single high-concentration exposure can destroy electrode completely.

- Prevention — Parylene-C barrier coating

| Advantages | Disadvantages |
| --- | --- |
| Physically blocks H₂O₂ from reaching silver | High-concentration H₂O₂ can still penetrate micropores over time |
| Protects against other liquid chemicals simultaneously | — |
| No change to electrode geometry | — |

- Prevention — User education (remove device before hair products)

| Advantages | Disadvantages |
| --- | --- |
| Zero cost, complete protection if followed | Relies entirely on user compliance |
| No impact on electrode design | Users forget — especially during routine habits |
| — | Poor user experience, friction in daily use |

5. Chloride regeneration from sweat

- Ag + Cl⁻ → AgCl + e⁻

- Electrode efficiency effect — positive Only beneficial reaction. Sweat chloride ions deposit fresh AgCl onto exposed silver surface. Self-heals minor AgCl layer losses. Impedance drops 30–50% after 10–15 minutes of wear. Signal quality improves with wear time, not degrades. Counteracts ammonia attack partially.

- No prevention needed — this works in your favour

6. Lactic acid from sweat

- AgCl + lactic acid → increased dissolution rate at low pH

- Electrode efficiency problem Slightly increases AgCl solubility at skin pH of 4.5–5.5. Minor and slow. Meaningful only in users with very acidic sweat or during intense exercise. Compounds ammonia attack — both dissolve AgCl simultaneously.

- Prevention — Parylene-C coating

| Advantages | Disadvantages |
| --- | --- |
| Slows acid penetration to AgCl surface | Least critical reaction — overengineering if used only for this |
| Same coating already used for other reactions — no extra step | Adds cost if not already using Parylene-C for other reasons |

7. Shampoo surfactant attack

- SLS + AgCl double layer → disrupts electrochemical interface

- Electrode efficiency problem Sodium lauryl sulfate disrupts the ionic double layer at electrode-skin interface. Impedance spikes from ~5kΩ to >80kΩ immediately after hair wash. Signal quality unusable for 30–60 minutes post-wash until skin naturally re-establishes ionic contact. Effect is temporary and fully recoverable.

- Prevention — Waterproof housing

| Advantages | Disadvantages |
| --- | --- |
| Complete protection from water and shampoo | Adds bulk to device design |
| User doesn't need to remove device to wash hair | Waterproof sealing increases manufacturing complexity |
| IPX4 or IPX7 rating achievable | Seals must be checked regularly for integrity |
| Only electrode contact face remains exposed | Trapped moisture inside housing if seal fails |

8. Silver ion release

- AgCl ⇌ Ag⁺ + Cl⁻   (Ksp = 1.8 × 10⁻¹⁰)

- Electrode efficiency effect — neutral to positive Release amount is negligible due to near-zero Ksp. No measurable impact on electrode performance. Ag⁺ ions at this concentration are clinically documented as antimicrobial — kills bacteria and fungi at skin contact zone. Prevents infection at behind-ear skin, especially relevant for long-term continuous wear.

- No prevention needed — net beneficial

- Combined prevention priority

- Most critical first:

- 1. Parylene-C coating      → covers reactions 1, 3, 4, 6, 7 simultaneously

- 2. Replaceable tip design  → covers reaction 3 long-term

- 3. Waterproof housing      → covers reaction 7 completely

- 4. Sealed storage case     → covers reaction 2 during non-use

- 5. User education          → covers reaction 4 as backup

- Parylene-C alone eliminates or slows five out of seven harmful reactions. Single highest-impact fabrication decision for your electrode.